## Slide 1: Chapter 4: Section IV

## Slide 2: Portable, Transport, and Non-invasive Ventilators
HAMILTON-T1
HAMILTON-MR1
Airon pNeuton
pNeuton mini
pNeuton A and S
Bio-Med Devices Crossvent 4+
Vyaire ReVel
Vyaire LTV 1200
Dräger CarinaDräger Oxylog 3000 Plus
Medtronic Newport HT 70 Plus
ZOLL Eagle II
Philips Respironics Trilogy Ventilator
DeVilbiss IntelliPAP Bilevel S
DeVilbiss IntelliPAP AutoBilevel
ResMed Lumis Tx
ResMed Astral 100/150

## Slide 3: Portable Ventilators HAMILTON-T1
Intended to provide positive pressure ventilatory support to adults and pediatrics, and optionally infants and neonates.21 It is pictured in (Figure 4-38). It is meant to be used in the ICU, intermediate care, emergency room, long-term acute care hospital or during transport.
An electronically controlled pneumatic ventilator with an integrated air compressing system.
It runs on AC or direct current (DC) power with battery backup.

## Slide 4: Portable Ventilators HAMILTON-T1 (cont’d)
Pneumatics deliver gas, and its electrical systems control pneumatics, monitor alarms, and distribute power. It employs a proximal flow sensor to monitor and deliver gas.
T1 operator interface utilizes a combination of numerics and real time waveforms and is depicted in (Figure 4-39).
It consists of a touch screen, some keys and a turn-and-press knob. Turning the knob changes the values of the parameter chosen and pressing the knob confirms the selection.

## Slide 5
Figure 4.38

## Slide 6
Figure 4.39

## Slide 8: Synchronized Controlled Mandatory Ventilation Plus [(S)CMV+]
Vt targeted, pressure-regulated mode. 21
Clinician sets the Vt and the ventilator uses varying levels of pressure to achieve the target.
The ventilator relies on calculations from previous breaths to deliver what it calculates the inspiratory pressure level needs to be to achieve the desired Vt.

## Slide 9: Main Operator Controls Include:
Vt target
PEEP
Fio2
I:E
Vt may not match the operator setting due to the ventilator is using inspiratory pressure to target the Vt as to VAC and relies on the previous breaths to make the calculation.21 Any change in patient effort from the previous breath may result in a deviation from the Vt setting.

## Slide 10: Synchronized, Intermittent Mandatory Ventilation Plus (SIMV+)
SIMV+ combines a Vt targeted PAC breath (see SCMV+ above) with spontaneous PS breaths in between.21
There exists mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a SCMV+ breath compounded by a PS breath) through the use of “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a SCMV+ breath.
If patient attempts to trigger a breath outside the window, then a spontaneous PS breath will occur.

## Slide 11: Main Operator Controls Include:
Vt target
PEEP
Fio2
Inspiratory time
PS level
Vt may not match operator setting for the SCMV+ breaths due to the ventilator is using inspiratory pressure to target the Vt as to VAC and relies on the previous breaths to make the calculation.21 Any change in patient effort from the previous breath may result in a deviation from the Vt setting.
During the SCMV+ breaths, the resulting pressure created is variable and depends on Vt, flow, and patient mechanics.
During PS breaths, the Vt will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.

## Slide 12: Pressure Controlled Ventilation+ (PCV+)
Pressure-targeted mode of ventilation that allows for mandatory, time-triggered breaths or patient-assisted breaths.21
Breath cycles to exhalation once the inspiratory time has been reached or when the high-pressure alarm is tripped.
Free breathing is allowed at both PEEP and inspiratory pressure levels.

## Slide 13: Main Operator Controls Include:
Inspiratory pressure level
I:E ratio
RR
PEEP
Fio2
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the inspiratory pressure level, patient effort, and patient respiratory system mechanics.21

## Slide 14: Pressure Synchronized, Intermittent Mandatory Ventilation+ (PSIMV+)
Similar to SIMV+ except the target is now pressure instead of volume and that inspiratory pressure is now set by the clinician.
P-SIMV+ also combines two breath types, PAC (PCV+) and PS or spontaneous breathing, while on CPAP.21
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through the use of a “timing window.”
If the patient attempts to trigger a breath in this timing window, then ventilator will deliver a PAC breath.
If the patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 15: Main Operator Controls Include:
Inspiratory pressure level
Inspiratory time
RR
PEEP
Fio2
Pressure support level
Cycle sensitivity for the PS breaths
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.21
During PS breaths, the volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.21

## Slide 16: Spontaneous (SPONT)
Patient triggers all breaths and will receive either PS or CPAP (if the PS is set at 0).21
If PS is set at 0, then the patient must contribute all of the work of breathing.
Sometimes used for a spontaneous breathing trial to assess the potential for liberation from the ventilator.
During PS, patient regulates the respiratory rate and the Vt with support from the ventilator.
Patient triggers all breaths and will receive the clinician-set PS level.
Breath cycles to expiration when the inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 17: Duo Positive Pressure Ventilation (DuoPAP)
Pressure-targeted mode of ventilation allowing for spontaneous breathing at two different pressure levels.21
Ventilator switches between two different clinician-set pressure levels (the lower of which can be CPAP).
Pressure levels are referred to as Phigh and PEEP. Thigh is the length of time at Phigh and Tlow is the length of time spent at Tlow.
Important difference from pressure-targeted modes is Phigh level is now referenced to atmosphere as opposed to the PEEP level.
Important: increases in Plow will decrease the distending pressure. Additional ventilator support can also be added in the form of PS at both pressure levels to augment the patient effort.

## Slide 18: Airway Pressure Release Ventilation (APRV)
Mode of ventilation similar to DuoPAP in allowing for spontaneous breathing at two different pressure levels.21
The ventilator switches between two different clinician-set pressure levels (the lower of which can be CPAP).
These pressure levels are referred to as Phigh and Plow. Thigh is length of time at Phigh and Tlow is length of time spent at Tlow.
Unlike DuoPAP, APRV generally is used in an inverse ratio fashion. Most of time spent at Phigh with brief, periodic releases to Plow.
Important difference from pressure-targeted modes is that the Phigh level is now referenced to atmosphere as to the PEEP level.
Important: increases in Plow will decrease distending pressure. Additional ventilator support can also be added in the form of PS at both pressure levels to augment the patient effort.

## Slide 19: Adaptive Support Ventilation (ASV)
“Feedback” mode that adjusts inspiratory pressure and RR to meet a minimum minute volume set by clinician.21
Results in a variable form of P-SIMV; machine-timed breaths are PAC while the spontaneously generated breaths are PS.
Amount of breathing that patient does determines how many machine-generated breaths will be delivered.
Algorithm for ASV is centered around lung protection.

## Slide 20: Noninvasive Ventilation
Patient triggers all breaths and will receive PS.21
NIV is designed for use with a mask or other noninvasive patient interface

## Slide 21: Spontaneous/Timed Noninvasive Ventilation (NIV-ST)
NIV-ST is also designed for use with a mask or other noninvasive patient interface.21
Essentially P-SIMV without an endotracheal tube. NIV-ST combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
If patient triggers regularly, all breaths are patient-triggered spontaneous breaths, i.e., no time-triggered mandatory breaths. Only when the patient trigger is not detected during defined breath cycle time (or total breath time), the ventilator delivers time-triggered mandatory breath.

## Slide 22: Spontaneous/Timed Noninvasive Ventilation (NIV-ST) (cont’d)
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in this timing window, then ventilator will deliver a PAC breath.
If the patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 23: Oxygen Special Features
High-pressure oxygen can be provided by a central gas supply or a gas cylinder (41 87 psi).21
Low-pressure oxygen can be provided by a concentrator or liquid cylinder (flow ≤ 15 L/min, ≤ 87 psi).

## Slide 24: Apnea Backup Ventilation
Ventilator will go into apnea backup ventilation if no breaths are detected in clinician-set interval.21
Mode of apnea backup ventilation depends on mode that patient is in (Table 4-51).

## Slide 25: Nebulizer
A pneumatic nebulization function powers a standard inline nebulizer for delivery of prescribed medications in ventilator circuit.21
When nebulization is active, nebulizer flow is synchronized with the inspiratory phase of each breath for 30 minutes.
Nebulization can be activated in all modes of ventilation.

## Slide 26: CO2 Monitoring
An optional CO2 sensor is available to allow monitoring of end tidal and volumetric expired CO2.21

## Slide 29: HAMILTON–MR1
HAMILTON-MR1 ventilator is intended to provide positive pressure ventilatory support to adults and pediatrics, and optionally infants and neonates.22 (pictured in Figure 4-40)
Intended for use in: the MRI department, the intensive care ward, intermediate care ward, emergency ward, long term acute care hospital or in the recovery room. MR1 can also be used as a transport ventilator.
Ventilator is classified as MR Conditional with the use of 1.5 Tesla and 3.0 Tesla static magnetic field scanners and is electronically controlled pneumatic ventilation system with an integrated air compressing system.

## Slide 30: HAMILTON–MR1 (cont’d)
It uses an external power supply connected to AC mains with battery backup. The ventilator’s pneumatics deliver gas, and its electrical systems control pneumatics, monitor alarms, and distribute power.
The user provides inputs to the microprocessor system through a touch screen, keys, and a press-and-turn knob then ventilator receives inputs from the proximal flow sensor and other sensors within the ventilator. (operator interface of the MR1 is pictured in Figure 4-41)

## Slide 31
Figure 4.40

## Slide 32
Figure 4.41

## Slide 34: Synchronized Controlled Mandatory Ventilation Plus ((S)CMV+)
Vt targeted, pressure-regulated mode.22
Clinician sets the Vt that want and ventilator uses varying levels of PAC to achieve the target.
Ventilator relies on calculations from previous breaths to deliver what it calculates the inspiratory pressure level needs to be to achieve the Vt.
In this mode, ventilator has the potential to drop inspiratory pressure level as patient effort increases and conversely may increase inspiratory pressure level when patient effort is too low.

## Slide 35: Main Operator Controls Include:
Fio2
PEEP
Vt target
I:E
RR
Vt may not match the operator setting due to the ventilator is using inspiratory pressure to target the Vt as to VAC and relies on the previous breaths to make the calculation.22 Any change in patient effort from previous breath may result in a deviation from the Vt setting.

## Slide 36: Synchronized, Intermittent Mandatory Ventilation Plus (SIMV+)
SIMV+ combines a Vt targeted PAC breath with spontaneous PS breaths in between.22
The type of breaths delivered are the same as P-SIMV; difference is the inspiratory pressure level is controlled by ventilator (variable) in SIMV+.

## Slide 37: Main Operator Controls Include:
Fio2
PEEP
Vt target
Inspiratory time
RR
PS level
Cycle sensitivity
Vt may not match operator setting for PAC breaths due to the ventilator is using inspiratory pressure to target the Vt as to VAC and relies on the previous breaths to make the calculation.22 Any change in patient effort from the previous breath may result in a deviation from Vt setting. During PS breaths, Vt will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.

## Slide 38: Pressure Controlled Ventilation+ PCV+
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.22
Spontaneous breaths not allowed.
Breath cycles to exhalation once the inspiratory time has been reached.

## Slide 39: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
I:E
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the inspiratory pressure level, patient effort, and patient respiratory system mechanics.22

## Slide 40: Pressure Synchronized, Intermittent Mandatory Ventilation+ (PSIMV+)
Similar to SIMV except the target is now pressure instead of volume.22
P-SIMV also combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through the use of “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 41: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure support level
Cycle sensitivity
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.22
During PS breaths, volume will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.22

## Slide 42: Spontaneous Mode
Patient triggers all breaths and will receive either PS or CPAP (if the PS is set at 0).22
If PS is set at 0 then patient must contribute all of the work of breathing.
Sometimes used for a spontaneous breathing trial to assess the potential for liberation from ventilator.
During Pressure Support, patient regulates the respiratory rate and Vt with support from ventilator.
Patient triggers all breaths and will receive clinician-set PS level.
Breath cycles to expiration when inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 43: Duo Positive Airway Pressure (DuoPAP)
Pressure-targeted mode of ventilation allowing for spontaneous breathing at two different pressure levels.22
The ventilator switches between two different clinician-set pressure levels (the lower of which can be CPAP).
These pressure levels are referred to as Phigh and Plow. Thigh is length of time at Phigh and Tlow is length of time spent at Tlow.
Important difference from pressure-targeted modes is the Phigh level is now referenced to atmosphere as to PEEP level.
Important: increases in Plow will decrease the distending pressure. Additional ventilator support can also be added in the form of PS at both pressure levels to augment the patient effort.

## Slide 44: Airway Pressure Release Ventilation (APRV)
Mode of ventilation similar to DuoPAP in allowing for spontaneous breathing at two different pressure levels.22
Ventilator switches between two different clinician-set pressure levels (the lower of which can be CPAP).
Pressure levels are referred to as Phigh and Plow. Thigh is length of time at Phigh and Tlow is length of time spent at Tlow Different from DuoPAP - APRV generally is used in an inverse ratio fashion.
Most of time is spent at Phigh with brief, periodic releases to Plow.
Important difference from pressure-targeted modes is that Phigh level is now referenced to atmosphere as to the PEEP level.
Important: increases in Plow will decrease the distending pressure. Additional ventilator support can also be added in the form of PS at both pressure levels to augment the patient effort.

## Slide 45: Adaptive Support Ventilation (ASV)
“Feedback” mode that adjusts the inspiratory pressure and RR to meet a minimum minute volume set by the clinician.22
This results in variable form of P-SIMV. The machine-timed breaths are PAC while the spontaneously generated breaths are PS.
Amount of breathing that patient does determines how many machine-generated breaths will be delivered.
Algorithm for ASV is centered around lung protection.

## Slide 46: Noninvasive Ventilation
In NIV mode, patient triggers all breaths and will receive either PS.22
NIV is designed for use with a mask or other noninvasive patient interface.

## Slide 47: Spontaneous/Timed Noninvasive Ventilation (NIV-ST)
Designed for use with a mask or other noninvasive patient interface.22 It is essentially P-SIMV without an endotracheal tube which combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
If patient triggers regularly, all breaths are patient-triggered spontaneous breaths, i.e., no time-triggered mandatory breath, only when patient trigger is not detected during the defined breath cycle time (or total breath time), ventilator delivers time-triggered mandatory breaths.

## Slide 48: Spontaneous/Timed Noninvasive Ventilation (NIV-ST) (cont’d)
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through the use of a “timing window.”
If patient attempts to trigger a breath in this timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 49: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure support level
Cycle sensitivity
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.22
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.22

## Slide 50: Nasal Continuous Positive Airway Pressure (nCPAP)
Available only in neonatal mode and is designed for use with a noninvasive interface.22
Purely spontaneous mode with CPAP as the baseline pressure with no additional pressure added on inspiration.

## Slide 51: Main Operator Controls Include:
Fio2
CPAP
RR
Vt created will be variable and depend on: patient effort, and patient respiratory system mechanics.22

## Slide 52: Nasal Continuous Positive Airway Pressure-pressure Control (nCPAP-PC)
Noninvasive pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.22
Breath cycles to exhalation once the inspiratory time has been reached

## Slide 53: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the inspiratory pressure level, patient effort, and patient respiratory system mechanics.22

## Slide 54: Apnea Backup Ventilation Special Features
Ventilator will go into apnea backup ventilation if no breaths are detected in the clinician-set interval.22
Mode of apnea backup ventilation depends on the mode that patient is in (table 4-37)

## Slide 55: Nebulizer
Pneumatic nebulization function powers a standard inline nebulizer for delivery of prescribed medications in ventilator circuit.22
When nebulization is active, nebulizer flow is synchronized with the inspiratory phase of each breath for 30 minutes.
Nebulization can be activated in all modes of ventilation.

## Slide 56: Dynamic Lung Panel
Dynamic lung panel visualizes tidal volume, lung compliance, patient triggering, and resistance in real time.22
Lungs expand and contract in synchrony with actual breaths. Numeric values for resistance (Rinsp) and compliance (Cstat) are displayed.
Shape of lungs and the bronchial tree are also related to compliance and resistance values.
If all values are in a normal range, the panel is framed in green.

## Slide 59: Accessories
HAMILTON-MR1 has an integrated vibrating mesh Aerogen® nebulizer for aerosolized medication delivery.22

## Slide 60: Airon pNeuton
Ventilators are small (< 15 pounds), pneumatically powered, pneumatically controlled and MRI compatible. The operator interfaces consist of a series of control knobs and a pressure gauge.
Being pneumatically driven they can provide ventilation while on transport or in the MRI suite.
pNeuton Mini is designed to ventilate neonatal, infant and pediatric patients.
pNeuton A and S models are designed to ventilate pediatric and adult patients. (The family of pNeuton ventilators are depicted in Figure 4-42)

## Slide 61
A. pNeuton Mini (neonatal/infant/pediatric)
B) pNeuton S (pediatric/adult)
C) pNeuton B (pediatric/adult)
Figure 4.42 A-C

## Slide 63: Continuous Positive Airway Pressure (CPAP) pNeuton Mini
Purely spontaneous mode of ventilation.23
No mandatory breaths are delivered.
Throughout breath cycle, clinician-set pressure is provided.
Breaths are supported only with demand flow.

## Slide 64: Main Operator Controls Include:
Fio2
CPAP
Vt will be variable and depend on patient effort, and patient respiratory system mechanics.23

## Slide 65: Intermittent Mandatory Ventilation (IMV) + CPAP
IMV delivers a set number of pressure-targeted (Pinsp) breaths in which inspiratory time is directly controlled by the clinician while the respiratory rate is controlled by expiratory time setting.23
In between these IMV breaths, patient can breathe spontaneously from baseline pressure setting (CPAP spontaneous).

## Slide 66: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
Expiratory time
Flow (Vinsp)

## Slide 67: Continuous Positive Airway Pressure pNeuton A and S
CPAP is a purely spontaneous mode of ventilation.24,25
No mandatory breaths are delivered.
Throughout the breath cycle the clinician-set pressure is provided.
Breaths are supported only with demand flow.

## Slide 68: Main Operator Controls Include:
Fio2
CPAP
Vt will be variable and depend on patient effort, and patient respiratory system mechanics.23

## Slide 69: Intermittent Mandatory Ventilation (IMV) + CPAP
IMV delivers a set number of pressure-targeted (Pinsp) breaths in which inspiratory time is directly controlled by clinician while respiratory rate is controlled by the expiratory time setting.23
In between these IMV breaths, the patient can breathe spontaneously from baseline pressure setting (CPAP spontaneous).

## Slide 70: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
Expiratory time
Flow (Vinsp)

## Slide 71: Continuous Positive Airway Pressure pNeuton A and S
CPAP is a purely spontaneous mode of ventilation.24,25
No mandatory breaths are delivered.
Throughout the breath cycle the clinician-set pressure is provided.
Breaths are supported only with demand flow.

## Slide 72: Main Operator Controls Include:
Fio2
CPAP
Vt will be variable and depend on patient effort, and patient respiratory system mechanics.23

## Slide 73: Intermittent Mandatory Ventilation (IMV) + CPAP with Pressure Limit
IMV delivers a set number of volume-targeted breaths in which the pressure limit is clinician-set.24,25
Inspiratory time and respiratory rate are directly controlled by the clinician.
Breath terminates either when the set tidal volume is delivered (volume target) or the pressure limit is met (pressure target).
Between these IMV breaths, patient can breathe spontaneously from the baseline pressure setting (CPAP spontaneous).

## Slide 74: Main Operator Controls Include:
Fio2
PEEP
Tidal volume
Peak pressure
Inspiratory time
Respiratory rate

## Slide 75
Special Features
Fio2 is reset at either 65% or 100%.24,25
All three ventilators can provide noninvasive ventilation.23-25
All three models are MRI compatible.23-25
Other
Ventilators are designed for patient transport.23-25
No leak compensation nor backup ventilation.23-25

## Slide 77: Bio-Med Devices Crossvent 4+
Bio-Med Devices actually provides a series of Crossvent ventilators: the CV 4+, CV 3+, CV 2+, and CV2i+.
CV 4+ is the most recent of the models and can ventilate all patient populations both invasively and noninvasively.
Ventilator is compact small (< 10 pounds), pneumatically powered, microprocessor controlled.
Operator interface consists three adjustment dials and a liquid crystal display. (The Bio-Med Crossvent 4+ ventilator is depicted in Figure 4-43)

## Slide 78
Figure 4.43

## Slide 80: Assist/Control (A/C)
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths or patient-assisted breaths.26
Spontaneous breathing not available.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered or when the maximum inspiratory pressure has been reached.

## Slide 81: Main Operator Controls Include:
PEEP
Vt
Inspiratory flow
RR
Max press
The resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.26

## Slide 82: Synchronized, Intermittent Mandatory Ventilation (SIMV)
Combines two breath types, VAC and either PS or spontaneous breathing, while on CPAP.26
There exists mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths will occur.

## Slide 83: Main Operator Controls Include:
PEEP
Vt
Inspiratory flow
SIMV rate
Pressure support level
During VAC breaths, resulting pressure created is variable and depends on Vt, flow, and patient mechanics.26
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.26

## Slide 84: Continuous Positive Airway Pressure (CPAP)
Purely spontaneous mode of ventilation.26 No mandatory breaths are delivered.
Throughout breath cycle, clinician-set pressure is provided. Pressure support can be added.
If a pressure support level > 0 cm H2O is set, breath cycles to expiration when inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 85: Main Operator Controls Include:
PEEP (CPAP)
Pressure support level
Vt will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.26 If the PS level is set to 0, all breaths will be unsupported.

## Slide 86
Special Features
Fio2 can be set at 50% or 100% with optional Air Entrainment Model, external blender can also be added.26
NIV is available but the mode(s) not specified.26
A backup rate can be set in CPAP in the event that apnea occurs.26
Other
Ventilators are designed for patient transport.26
No leak compensation nor nebulizer available.26

## Slide 88: Vyaire ReVel
The Vyaire ReVel ventilator is designed for pediatric (> 5 kg) and adult patients during transport or in subacute care situations.27 It is pneumatically driven and electronically controlled. (pictured in Figure 4-44)
Operator interface consists of a series of knobs and LED readouts (depicted in Figure 4-45)
The ventilator comes with a docking station which can interface with a graphics monitor and nurse call system.

## Slide 89
Figure 4.44

## Slide 90
Figure 4.45

## Slide 92: A/C + Volume
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths or patient-assisted breaths.27
Spontaneous breathing not allowed.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered.

## Slide 93: Main Operator Controls Include:
Fio2
PEEP
Vt
Inspiratory flow
RR
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.27

## Slide 94: Synchronized, Intermittent Mandatory Ventilation + Volume - SIMV + Volume
Combines two breath types, VAC and either PS or spontaneous breathing, while on CPAP.27
Mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths are all provided.
The ventilator prevents “breath stacking” (a VAC breath compounded by a PS breath) through use of “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a VAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 95: Main Operator Controls Include:
Fio2
PEEP
Vt
Flow
SIMV rate
Pressure support level
Cycle sensitivity
During VAC breaths, resulting pressure created is variable and depends on Vt, flow, and patient mechanics.27
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.27

## Slide 96: A/C + Pressure
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.27
Spontaneous breaths not allowed.
Breath cycles to exhalation once the inspiratory time has been reached or when the high-pressure alarm is tripped.

## Slide 97: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the inspiratory pressure level, patient effort, and patient respiratory system mechanics.27

## Slide 98: SIMV + Pressure
Similar to SIMV + Volume except the target is now pressure instead of volume.27
Combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in this timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 99: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
SIMV rate
Pressure support level
Cycle sensitivity
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.27
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.27

## Slide 100: CPAP + Pressure Support (CPAP + PSV)
Patient regulates respiratory rate and Vt with support from the ventilator.27
Patient triggers all breaths and will receive the clinician-set PS level.
If a pressure support level > 0 cm H2O is set, the breath cycles to expiration when the inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 101: Main Operator Controls Include:
Fio2
PEEP
Pressure support level
Cycle sensitivity
Vt will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.27 If the PS level is set to 0, all breaths will be unsupported.

## Slide 102: A/C + Pressure-regulated Volume Control (A/C + PRVC)
Pressure-targeted (either PC or PA breaths) mode that uses pressure breaths to achieve a clinician-set Vt.27
The inspiratory pressure level; however, is now adjusted by the ventilator instead of the clinician.
Ventilator relies a previous breath sequence to deliver the ensuing breath the Vt tends to be somewhat variable, especially in patients with inconsistent efforts.
No spontaneous breaths.

## Slide 103: Main Operator Controls Include:
Fio2
PEEP
RR
Vt target
Inspiratory time
Amount of pressure and flow delivered will be variable and depend on the Vt setting, patient effort, and patient respiratory system mechanics.27

## Slide 104: SIMV + PRVC
Combines two breath types, PAC (ventilator-controlled inspiratory pressure) and either PS or spontaneous breathing, while on CPAP.27
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in this timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 105: Main Operator Controls Include:
Fio2
PEEP
SIMV rate
Vt
Inspiratory time
Pressure support level
Cycle sensitivity
During control breath amount of pressure and flow delivered will be variable and depend on the Vt setting, patient effort, and patient respiratory system.27 During PS breath flow, the Vt will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics. If CPAP is used, the spontaneous breath Vt will be variable and depend on patient effort, and patient respiratory system mechanics.

## Slide 106: Pressure-regulated Volume Support (VS)
Ventilator delivers a variable (from breath to breath) PS level in order to reach a clinician-set Vt.27
Ventilator changes PS level as needed to reach Vt.
Ventilator drops the level of support as patient effort increases and conversely increases PS level when patient effort is too low.

## Slide 107: Main Operator Controls Include:
Fio2
PEEP
Vt target
Cycle sensitivity
Vt may not match operator setting due to the ventilator is using PS to target the Vt as to VAC and relies on previous breaths to make calculation.27 Any change in patient effort from previous breath may result in a deviation from Vt setting.

## Slide 108: Noninvasive Positive Pressure Ventilation A/C (NPPV A/C)
A/C + Pressure (see above) delivered through a mask instead of an endotracheal tube.27
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.
Spontaneous breaths not allowed.
Breath cycles to exhalation once the inspiratory time has been reached or when the high-pressure alarm is tripped.

## Slide 109: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the inspiratory pressure level, patient effort, and patient respiratory system mechanics.27

## Slide 110: Noninvasive Positive Pressure Ventilation/CPAP/Pressure Support (NPPV/CPAP/PS)
Is CPAP + PSV delivered through a mask instead of an endotracheal tube.27
During NPPV/CPAP/PSV the patient regulates, respiratory rate and Vt with support from ventilator.
Patient triggers all breaths and will receive the clinician-set PS level.
If pressure support level > 0 cm H2O is set, breath cycles to expiration when the inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 111: Main Operator Controls Include:
Fio2
PEEP
Pressure support level
Cycle sensitivity
Vt will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.27 If PS level is set to 0, all breaths will be unsupported.

## Slide 112: Apnea Backup Ventilation Special Features
The ventilator engages apnea backup ventilation when no breath has been delivered by the time clinician-selected apnea interval elapses.27
When in A/C or SIMV modes, apnea backup rate is determined by the operator-set mandatory breath rate or the apnea interval setting (whichever provides the highest respiratory rate).
In CPAP + PS apnea backup ventilation is provided in PAC.
When apnea interval setting (found in the alarm limits window) determines the backup rate, ventilator will continue to ventilate at this rate until the apnea has been resolved.
All other controls for apnea ventilation in A/C and SIMV are set when primary control values for these modes are selected.

## Slide 113: Oxygen
When ventilator is attached to a high-pressure oxygen source of 40 psi to 88 psi the percentage of oxygen to be delivered through the ventilator’s oxygen blending system can be set directly.27
When O2% control is set to low-pressure source (LPS), oxygen can be supplied from a low-pressure, low-flow oxygen source of < 10 psi such as a flow meter.
Oxygen from low-pressure source is mixed with air inside the ventilator once connected to a low-pressure source, the estimated O2 percent delivered to patient is determined by the O2 inlet flow and the minute volume and is not regulated by ventilator.
A separate chart is provide to help estimate the delivered Fio2.

## Slide 114: Spontaneous Breathing Trial
During a spontaneous breathing trial, various parameters can be displayed.27
These include SBT f/Vt, SBT respiratory rate, SBT Vt and SBT time remaining.

## Slide 115: Nebulizer
Nebulization only available during Volume breaths in Assist/Control mode only.27
When Nebulizer is activated, a six (6) L/min nominal flow is delivered to the nebulizer drive port.
This drives an aerosol nebulizer that doses medication into patient circuit.

## Slide 116: Other
Optional pulse oximetry is available along with a separate graphics monitor.27

## Slide 119: Vyaire LTV 1200
Designed for pediatric (> 11 lbs) and adult patients in institutional, home, and transport settings.28
Blower driven and electronically controlled and can ventilate either invasively of noninvasively. (pictured in Figure 4-46)
The operator interface consists of a series of buttons, LED readouts, and an adjustable control knob (Figure 4-47).

## Slide 120
Figure 4.46

## Slide 121
Figure 4.47

## Slide 123: Volume A/C
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths or patient-assisted breaths.28
Spontaneous breathing not allowed.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered.

## Slide 124: Main Operator Controls Include:
Fio2
PEEP
Vt
Inspiratory flow
RR
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.28

## Slide 125: Volume SIMV
Volume SIMV combines two breath types, VAC and either PS or spontaneous breathing, while on CPAP.28
Mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths are all provided.
Ventilator prevents “breath stacking” (a VAC breath compounded by a PS breath) through the use of a “timing window.”
If patient attempts to trigger a breath in this timing window, then ventilator will deliver a VAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 126: Main Operator Controls Include:
Fio2
PEEP
Vt
Flow
SIMV rate
Pressure support level
Cycle sensitivity for the PS breaths
During VAC breaths, resulting pressure created is variable and depends on Vt, flow, and patient mechanics.28
During PS breaths, volume will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.28

## Slide 127: Pressure A/C
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.28
Spontaneous breaths not allowed.
Breath cycles to exhalation once the inspiratory time has been reached or when the high-pressure alarm is tripped.

## Slide 128: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the inspiratory pressure level, patient effort, and patient respiratory system mechanics.28

## Slide 129: Pressure SIMV
Similar to Volume SIMV except the target is now pressure instead of volume.28 P-SIMV also combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 130: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
SIMV rate
Pressure support level
Cycle sensitivity for the PS breaths
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.28
During PS breaths, volume will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.28

## Slide 131: CPAP/Pressure Support
Patient regulates respiratory rate and the Vt with support from ventilator.28
Patient triggers all breaths and will receive the clinician-set PS level.
If a pressure support level > 0 cm H2O is set, the breath cycles to expiration when the inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 132: Main Operator Controls Include:
Fio2
PEEP
Pressure support level
Cycle sensitivity
Vt will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.28 If the PS level is set to 0, all breaths will be unsupported

## Slide 133: Noninvasive Positive Pressure Ventilation
NPPV is PS delivered through a mask instead of an endotracheal tube.28
Patient regulates the respiratory rate and the Vt with support from ventilator.
Patient triggers all breaths and will receive clinician-set PS level.
If a pressure support level > 0 cm H2O is set, the breath cycles to expiration when the inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 134: Main Operator Controls Include:
Fio2
PEEP
Pressure support level
Cycle sensitivity
Vt will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.28 If the PS level is set to 0, all breaths will be unsupported.

## Slide 135: Apnea Backup Ventilation Special Features
Ventilator engages Apnea Backup Ventilation when no breath has been delivered by the time the clinician-selected apnea interval elapses.28
Backup ventilation will occur in the assist/control mode related to current settings (Table 4-43).

## Slide 136: Oxygen
Ventilator is attached to a high-pressure oxygen source of 40 psi to 80 psi.28
Percentage of oxygen to be delivered through ventilator’s oxygen blending system can be set directly.
When O2% control is set to low-pressure source (LPS), oxygen can be supplied from a low-pressure, low-flow oxygen source of < 35 psi such as a flow meter.
Oxygen from the low-pressure source is mixed with air inside the ventilator.
When connected to a low-pressure source, the estimated O2 percent delivered to patient is determined by O2 inlet flow and the minute volume and is not regulated by ventilator.
A separate chart is provide to help estimate delivered Fio2.

## Slide 137: Spontaneous Breathing Trial
During a spontaneous breathing trial various parameters can be displayed.28
Including SBT f/Vt, SBT respiratory rate, SBT Vt and SBT time remaining.

## Slide 138: Nebulizer
Only available during Volume breaths in Assist/Control mode only.28
When Nebulizer is activated, a six (6) L/min nominal flow is delivered to the nebulizer drive port.
This drives an aerosol nebulizer that doses medication into patient circuit.

## Slide 139: Other
There is no leak compensation nor nebulizer in the LTV 1200.28

## Slide 142: Dräger Carina
Designed for treatment of sub-acute care patients in hospitals or medical rooms and can be used for transport and in the hospital setting.29
Lightweight ventilator (5.5 kg) that is pneumatically powered through blender technology and electrically driven.
Carina can deliver invasive or noninvasive ventilation and is (pictured in Figure 4-48)
Operator interface consists of an LCD screen with an array of buttons and a control knob.

## Slide 143
Figure 4.48

## Slide 145: Volume Control Synchronized, Intermittent Mandatory Ventilation with Autoflow (VC-SIMV)
Combines two breath types, PAC (ventilator-controlled inspiratory pressure) and either PS or spontaneous breathing, while on CPAP.29
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in this timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 146: Main Operator Controls Include:
Fio2
PEEP
SIMV rate
Vt
Inspiratory time or I:E
Pressure support level
During control breath amount of pressure and flow delivered will be variable and depend on the Vt setting, patient effort, and patient respiratory system mechanics.29
During PS breath flow, Vt will be variable and depend on PS level, patient effort, and patient respiratory system mechanics. If CPAP is used, spontaneous breath Vt will be variable and depend on patient effort, and patient respiratory system mechanics.

## Slide 147: Pressure Control-A/C (PC-A/C)
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.29
Spontaneous breaths not allowed.
Breath cycles to exhalation once inspiratory time has been reached or when the high-pressure alarm is tripped.

## Slide 148: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time or I:E
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the inspiratory pressure level, patient effort, and patient respiratory system mechanics.29

## Slide 149: Pressure Control-biphasic Positive Airway Pressure (PC-BIPAP)
Form of pressure SIMV that combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.29
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Spontaneous breathing allowed throughout entire breath cycle.

## Slide 150: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time or I:E
SIMV rate
Pressure support level
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.29
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.

## Slide 151: Spontaneous – Continuous Positive Airway Pressure/Pressure Support (SPN-CPAP/PS)
Patient regulates the respiratory rate and Vt with support from ventilator.29
Patient triggers all breaths and will receive clinician-set PS level.
Breath cycles to expiration when inspiratory flow decreases below a preset percentage of inspiratory peak flow.

## Slide 152: Main Operator Controls Include:
PEEP
Fio2
Pressure support level
Vt will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.29

## Slide 153: Spontaneous - Continuous Positive Airway Pressure (SPN-CPAP)
Ventilator maintains a baseline pressure level (CPAP) and patient breaths spontaneously.29
Patient receives no additional support from ventilator.

## Slide 154: Main Operator Controls Include:
Fio2
CPAP
Vt may not match the operator setting due to the ventilator is using PS to target the Vt as to VAC and relies on previous breaths to make the calculation.29 Any change in patient effort from the previous breath may result in a deviation from Vt setting.

## Slide 155: Spontaneous–continuous Positive Airway Pressure/Pressure Support with Volume Guarantee (SPN-CPAP/PS)
Guarantee the ventilator delivers a variable (from breath to breath) PS level in order to reach a clinician-set Vt.29
Ventilator changes the PS level as needed to reach Vt.
Ventilator drops the level of support as patient effort increases and conversely increases PS level when patient effort is too low.

## Slide 156: Main Operator Controls Include:
Fio2
PEEP
Vt target
Vt may not match the operator setting due to the ventilator is using PS to target the Vt as to VAC and relies on previous breaths to make the calculation.29 Any change in patient effort from previous breath may result in a deviation from Vt setting.

## Slide 157: Backup Ventilation Special Features
Available in SPN-CPAP, SPN-CPAP/PS and SPN-CPAP/PS with volume guarantee.29
Ventilator will deliver VC-SIMV with AutoFlow breaths.

## Slide 158: Oxygen
Oxygen levels of 21% to 100% are delivered when supplied with a high pressure of 40 to 87 psi.29
Low-pressure (> 7 psi) oxygen delivery is also available.
Delivered Fio2 will vary depending on the O2 inlet flow and the minute volume.

## Slide 159: Noninvasive Ventilation
Available in all available modes.29
Ventilator will adjust the trigger level based on the leak when NIV is selected.

## Slide 160: Other
Nebulizer and neonatal ventilation not available.29

## Slide 162: Dräger Oxylog 3000 Plus
Time-cycled, volume-controlled and pressure-controlled emergency and transport ventilator for patients requiring mandatory or assisted ventilation with a tidal volume from 50 mL upwards.30
Oxylog 3000 Plus can deliver invasive or noninvasive ventilation (pictured in Figure 4-49)
Operator interface consists of an LCD screen with an array of buttons and a control knob. (pictured in Figure 4-50)

## Slide 163
Figure 4.49

## Slide 164
Figure 4.50

## Slide 166: Volume Control–continuous Mandatory Ventilation (VC-CMV)
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths only.30
Patient assisted breaths, and spontaneous breathing not allowed.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered.

## Slide 167: Main Operator Controls Include:
PEEP
Fio2
Vt
Inspiratory flow
RR
Inspiratory time
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.30

## Slide 168: Volume Control–A/C (VC–A/C)
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths or Patient assisted breaths.30
Spontaneous breathing not allowed.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered.

## Slide 169: Main Operator Controls Include:
PEEP
Fio2
Vt
Inspiratory flow
RR
Inspiratory time
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.30

## Slide 170: Volume Control Synchronized, Intermittent Mandatory Ventilation (VC-SIMV)
Combines two breath types, VAC and either PS or spontaneous breathing, while on CPAP.30
Mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths are all provided.
Ventilator prevents “breath stacking” (a VAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in this timing window, then ventilator will deliver a VAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 171: Main Operator Controls Include:
PEEP
Fio2
Vt
Flow
SIMV rate
Inspiratory time
Pressure support level
Cycle sensitivity for the PS breaths
During VAC breaths, resulting pressure created is variable and depends on Vt, flow, and patient mechanics.30
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.30

## Slide 172: Volume Control–continuous Mandatory Ventilation with Autoflow (VC-CMV)
Pressure-targeted (PC breaths) mode that uses pressure breaths to achieve a clinician-set Vt.30
Inspiratory pressure level is now adjusted by the ventilator instead of clinician.
Ventilator relies a previous breath sequence to deliver ensuing breath the Vt tends to be somewhat variable, especially in patients with inconsistent efforts.
This mode delivers mandatory, time-triggered breaths only.
No spontaneous breaths allowed.
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths only. Patient assisted breaths, and spontaneous breathing not allowed.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered.

## Slide 173: Main Operator Controls Include:
PEEP
Fio2
Vt target
Inspiratory flow
RR
Inspiratory time
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.30

## Slide 174: Volume Control–A/C with Autoflow (VC-A/C)
Pressure-targeted (either PC or PA breaths) mode that uses pressure breaths to achieve a clinician-set Vt.30
Inspiratory pressure level, is now adjusted by the ventilator instead of clinician.
Since ventilator relies a previous breath sequence to deliver the ensuing breath the Vt tends to be somewhat variable, especially in patients with inconsistent efforts.
No spontaneous breaths allowed.

## Slide 175: Main Operator Controls Include:
Fio2
PEEP
RR
Vt target
Inspiratory time
Amount of pressure and flow delivered will be variable and depend on the Vt setting, patient effort, and patient respiratory system mechanics.30

## Slide 176: Volume Control Synchronized, Intermittent Mandatory Ventilation with Autoflow (VC-SIMV)
Combines two breath types, PAC (ventilator-controlled inspiratory pressure) and either PS or spontaneous breathing, while on CPAP.30
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 177: Main Operator Controls Include:
Fio2
PEEP
SIMV rate
Vt
Inspiratory time
Pressure support level
During control breath, amount of pressure and flow delivered will be variable and depend on Vt setting, patient effort, and patient respiratory system mechanics.30 During PS breath flow, the Vt will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics. If CPAP is used, the spontaneous breath Vt will be variable and depend on patient effort, and patient respiratory system mechanics.

## Slide 178: Pressure Control-biphasic Positive Airway Pressure (PC-BIPAP)
Form of pressure SIMV that combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.30
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Spontaneous breathing allowed throughout entire breath cycle.

## Slide 179: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
SIMV rate
Pressure support level
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.30
During PS breaths, volume will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.30

## Slide 180: Spontaneous-continuous Positive Pressure (SPN-CPAP)
Patient regulates the respiratory rate and Vt with support from ventilator.30
Patient triggers all breaths and will receive clinician-set PS level or no additional support (CPAP).
If a pressure support level > 0 cm H2O is set, the breath cycles to expiration when inspiratory flow decreases below a preset percentage of inspiratory peak flow.

## Slide 181: Main Operator Controls Include:
PEEP
Fio2
Pressure support level
Vt will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.30 If the PS level is set to 0, all breaths will be unsupported.

## Slide 182: Backup Ventilation Special Features
Available in SPN-CPAP only.30
Ventilator will deliver VC-CMV breaths.

## Slide 183: Oxygen
40% to 100% when supplied with a high pressure of 39 to 87 psi through either a medical pipeline system or portable oxygen tank.30
Actual Fio2 may vary somewhat depending on inspiratory flowrate and mean airway pressure.

## Slide 184: Noninvasive Ventilation
Available in SpnCPAP (/PS), PC-BIPAP, VC-CMV /AF, VC-A/C /AF and VC-SIMV/AF.30
Ventilator will compensate for inherent leaks and the leakage alarm will be deactivated.

## Slide 185: Other
An optional CO2 sensor is available to allow for measurement of exhaled CO2.30
Nebulizer and neonatal ventilation not available.30
Previous versions of this ventilator include the Oxylog 3000, Oxylog 2000 Plus and the Oxylog 1000.30
Internal battery will operate for approximately 7.5 to 9.5 hours (depending on the parameter settings) on full charge.30

## Slide 187: Medtronic Newport HT 70 Plus
Ventilator is designed to provide ventilatory support for infant, pediatric and adult patients in emergency care, transport, critical care, subacute care, home care and emergency preparedness applications.31
Electronically controlled and functions with a twin micro-piston pump to power the pneumatic system and has servo-controlled, leak-compensated PEEP.
HT70 may be used with an endotracheal tube, tracheal tube, facemask, nasal mask or prongs, or a mouthpiece. (pictured in Figure4-51)
Operator interface uses a color-coded LCD touch screen featuring an array of buttons along with a display screen. Settings are entered by selecting a parameter and then adjusting the value using the up/down arrow buttons. (pictured in Figure 4-52)

## Slide 188
Figure 4.51

## Slide 189
Figure 4.52

## Slide 191: Assist/Control Mandatory Ventilation Volume Control (A/CMV VC)
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths or patient-assisted breaths.31
Spontaneous breathing not allowed.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered.

## Slide 192: Main Operator Controls Include:
Fio2
PEEP
Vt
Inspiratory flow
Inspiratory time (Tinsp)
RR
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.31

## Slide 193: Synchronized, Intermittent Mandatory Ventilation Volume Control (SIMV VC)
Combines two breath types, VAC and either PS or spontaneous breathing, while on CPAP.31
Mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths are all provided.
Ventilator prevents “breath stacking” (a VAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a VAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 194: Main Operator Controls Include:
Fio2
PEEP
Vt
Inspiratory flow
Inspiratory time
SIMV rate
Pressure support level
Cycle sensitivity
During VAC breaths, resulting pressure created is variable and depends on Vt, flow, and patient mechanics.31
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.31

## Slide 195: Assist/Control Mandatory Ventilation Pressure Control (A/CMV PC)
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.31
Spontaneous breaths not allowed.
Breath cycles to exhalation once the inspiratory time has been reached or when the high-pressure alarm is tripped.

## Slide 196: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the inspiratory pressure level, patient effort, and patient respiratory system mechanics.31

## Slide 197: Synchronized, Intermittent Mandatory Ventilation Pressure Control (SIMV PC)
Similar to SIMV/VC except the target is now pressure instead of volume.31
SIMV/PC also combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through the use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 198: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
SIMV rate
Pressure support level
Cycle sensitivity for the PS breaths
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.31
During PS breaths, volume will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.31

## Slide 199: Spontaneous Ventilation (SPONT Mode)
Patient regulates respiratory rate and the Vt.31
Patient triggers all breaths and will receive either CPAP or a clinician-set PS level.
If a pressure support level > 0 cm H2O is set, the breath cycles to expiration when the inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 200: Main Operator Controls Include:
Fio2
PEEP
Pressure support level
Cycle sensitivity
Vt will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.31 If the PS level is set to 0, all breaths will be unsupported.

## Slide 201: Backup Ventilation Special Features
Occurs when either, low minute volume (MVI/MVE) or apnea alarm, or both, are activated.31
Available in all of ventilator modes.
If current mode is A/CMV or SIMV, current settings are used with a respiratory rate of 1.5 times the current setting.
If current mode is SPONT Mode, ventilator delivers SIMV PC breaths at a RR of 15, inspiratory pressure 15 cm H2O above set PEEP and a 1.0 second inspiratory time.

## Slide 202: Oxygen
HT 70 Plus can use a 35 to 65 psig source and can deliver an Fio2 from 0.21 to 1.0.31

## Slide 203: Battery Backup
Internal Dual Battery System consists of two internal independent but coordinated lithium ion batteries, one located on the back of the ventilator and the secondary backup battery inside the ventilator.31
Internal Dual Battery System can provide up to 10 hours of operation at standard settings when new and fully charged.

## Slide 204: NIV
HT70 Plus can be used for noninvasive ventilation in all modes.31 It is activated switching the NIV toggle switch to on.
When NIV is activated the bias flow is increased to 10 L/min (and can be adjusted from 3-30 L/min), the low minute volume alarm can be turned off, the low-pressure alarm can be set closer to the base pressure (1 cm H2O above baseline), and the high minute volume alarm range is expanded to 80 L/min.

## Slide 207: ZOLL Eagle II
A full-feature, portable, mechanical ventilator designed to operate in hospitals, during transport, and in severe environments, where it may be exposed to rain, dust, rough handling, and extremes in temperature and humidity (Figure 4-53).32
With a third-party filter in place, ventilator may be operated in environments where chemical and/or biological toxins are present. When marked with an “MRI conditional” label, they are suitable for use in an MRI environment with appropriate precautions.
The Eagle II is a volume and pressure-targeted, time- or flow-cycled ventilator designed to use either oxygen from a 55 psig source or fresh air using its internal compressor to deliver a positive-pressure breath. When Fio2 is set to 21%, clinicians can connect a low-flow oxygen source at the Fresh Gas/Emergency Air Intake using a reservoir.

## Slide 208: ZOLL Eagle II (cont’d)
Operator interface (Figure 4-54) consists of a liquid crystal display (LCD) provides continuous display of control settings, operating conditions, power, and alarm status information. Most unit functions are controlled by pressing the parameter button associated with parameter you wish to change.
Pressing parameter button highlights the primary parameter; additional presses highlight secondary parameters moving in a clockwise direction.
When parameter you wish to change is highlighted, a knob called the rotary encoder is used to adjust parameter to the desired setting, and then it is confirmed by pressing confirm/select button.

## Slide 209
Figure 4.53

## Slide 210
Figure 4.54

## Slide 212: Assist/Control Volume Target
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths or patient-assisted breaths.32
Spontaneous breathing not allowed.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered. Inspiratory flow is preset to be a square waveform.

## Slide 213: Main Operator Controls Include:
Fio2
PEEP
Vt
I:E or inspiratory time
RR
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.32

## Slide 214: Synchronized, Intermittent Mandatory Ventilation (SIMV) Volume Target
Combines two breath types, VAC and either PS or spontaneous breathing, while on CPAP.32
Mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths are all provided.
Ventilator prevents “breath stacking” (a VAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a VAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur. Inspiratory flow is preset to be a square waveform.

## Slide 215: Main Operator Controls Include:
Fio2
PEEP
Vt
I:E or inspiratory time
SIMV rate
Pressure support level
Cycle sensitivity
During VAC breaths, resulting pressure created is variable and depends on Vt, flow, and patient mechanics.32
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.32

## Slide 216: Assist/Control (A/C) Pressure Target
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.32
Spontaneous breaths not allowed.
Breath cycles to exhalation once the inspiratory time has been reached or when high-pressure alarm is tripped.

## Slide 217: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
I:E or inspiratory time
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on inspiratory pressure level, patient effort, and patient respiratory system mechanics.32

## Slide 218: Synchronized, Intermittent Mandatory Ventilation (SIMV) Pressure Target
Similar to SIMV Volume Target described above except the target is now pressure instead of volume.32
SIMV Pressure Target also combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in this timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 219: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
I:E or inspiratory time
SIMV rate
Pressure support level
Cycle sensitivity
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.32
During PS breaths, volume will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.32

## Slide 220: Continuous Positive Airway Pressure (CPAP)
Patient regulates respiratory rate and the Vt with support from ventilator.32
Patient triggers all breaths and will receive either a purely spontaneous breath (demand flow only available) or the clinician-set PS.
If a pressure support level > 0 cm H2O is set, breath cycles to expiration when the inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 221: Main Operator Controls Include:
Fio2
PEEP
Pressure support level
Cycle sensitivity

## Slide 222: Noninvasive Positive Pressure Ventilation (NPPV)
Provides flow during expiratory phase to maintain baseline pressure (CPAP) in spontaneously breathing patients with a leaking airway or facemask.32
Amount of leak compensation depends on leak flow rate during expiratory period and ranges from 0 to 15 liters/min and is automatically adjusted by the ventilator in order to maintain the CPAP target.
NPPV/CPAP/PS is CPAP/PS (see above) delivered through a mask instead of an endotracheal tube. During NPPV, patient regulates respiratory rate and Vt.

## Slide 223: Main Operator Controls Include:
Fio2
CPAP
Vt will be variable and depend on patient effort, and patient respiratory system mechanics.32

## Slide 224: Oxygen Special Features
The unit can use oxygen from low-flow sources, oxygen flow meters, and oxygen concentrators to provide supplemental oxygen to patients.32
To do this, oxygen is delivered through Fresh Gas/Emergency Air Intake when unit’s internal compressor cycles to deliver a breath.

## Slide 225: Backup Ventilation
Depending upon preexisting conditions at time of failure, backup ventilator will begin operation in one of two ways:
  (1) If no preexisting alarm condition(s) exists, backup operation will continue using the current settings.
  (2) If a preexisting alarm condition(s) exists, backup operation will revert to the startup default settings (Mode A/C (P), volume target, BPM 12, PIP 20 cm H2O, Fio2 21%, PEEP 5 cm H2O, I:E 1:2.5, PIP High Limit 35 cm H2O).32

## Slide 226: Other
An optional pulse oximeter can be used.32

## Slide 228: Philips Respironics Trilogy Ventilator
Respironics Trilogy 202 systems provides continuous or intermittent ventilatory support for the care of individuals who require mechanical ventilation with or without Fio2 blending.33 (pictured in Figure 4-55) and intended for pediatric through adult patients weighing at least 5 kg.
The ventilator is meant to be used in hospitals and institutions, and for portable applications such as wheelchairs and gurneys only when in an institutional setting and may be used for both invasive and noninvasive ventilation.
Not intended to be used as a transport ventilator.
The Trilogy 202 can be used with either a single-limb or double-limb circuit and operator interface consists of a central screen with some buttons to make parameter and setting adjustments and is (depicted in Figure 4-56)

## Slide 229
Figure 4.55

## Slide 230
Figure 4.56

## Slide 232: Control Ventilation (CV)
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths only.33
Patient triggered breaths, and spontaneous breathing are not allowed.
Breath cycles to exhalation once tidal volume (Vt) has been delivered.

## Slide 233: Main Operator Controls Include:
Fio2
PEEP
Vt
Inspiratory time
RR
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.33

## Slide 234: Assist/Control (A/C)
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths only or patient-assisted breaths.33
Breath cycles to exhalation once the tidal volume (Vt) has been delivered.

## Slide 235: Main Operator Controls Include:
Fio2
PEEP
Vt
Inspiratory time
RR
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.33

## Slide 236: Synchronized, Intermittent Mandatory Ventilation (SIMV)
Combines two breath types, VAC and either PS or spontaneous breathing, while on CPAP.33
Mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths are all provided.
Ventilator prevents “breath stacking” (a VAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a VAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 237: Main Operator Controls Include:
Fio2
PEEP
Vt
Inspiratory time
SIMV rate
Pressure support level
Cycle sensitivity
During VAC breaths, resulting pressure created is variable and depends on Vt, flow, and patient mechanics.33
During PS breaths, volume will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.33

## Slide 238: Pressure-control Mode (PC)
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.33
Spontaneous breaths not allowed.
Breath cycles to exhalation once the inspiratory time has been reached.

## Slide 239: Main Operator Controls Include:
Fio2
EPAP (expiratory positive airway pressure)
IPAP (inspiratory positive airway pressure)
I-Time
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on the IPAP, EPAP levels, patient effort, and patient respiratory system mechanics.33

## Slide 240: Pressure Control–synchronized, Intermittent Mandatory Ventilation (PC-SIMV)
Similar to SIMV except the target is now pressure instead of volume.33
Combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through the use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 241: Main Operator Controls Include:
Fio2
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure support level
Cycle sensitivity
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.33
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.33

## Slide 242: Continuous Positive Airway Pressure (CPAP)
Ventilator functions as a demand flow system, with the patient triggering all breaths and determining their timing, pressure, and size.33

## Slide 243: Main Operator Controls Include:
Fio2
CPAP level
Vt will be variable and depend on patient effort, and patient respiratory system mechanics.33

## Slide 244: Spontaneous Mode
Delivers pressure-targeted breaths that are patient-triggered and flow-cycled.33

## Slide 245: Main Operator Controls Include:
Fio2
IPAP
EPAP
Flow (expiratory) sensitivity
Vt will be variable and depend on IPAP, EPAP, patient effort, and patient respiratory system mechanics.33

## Slide 246: Spontaneous/Timed (S/T)
Delivers pressure-controlled, time-cycled mandatory and pressure supported spontaneous breaths.33
If patient fails to trigger a breath within the interval determined by the rate setting, ventilator triggers a mandatory breath with the set I-Time.

## Slide 247: Main Operator Controls Include:
Fio2
EPAP
IPAP
I-Time
RR
Vt will be variable and depend on IPAP, EPAP, patient effort, and patient respiratory system mechanics.33

## Slide 248: Timed (T Mode)
Delivers pressure-controlled, time-cycled mandatory breaths only.33
Patient-triggered and spontaneous breaths not allowed.

## Slide 249: Main Operator Controls Include:
Fio2
EPAP
IPAP
I-Time
RR
Vt will be variable and depend on IPAP, EPAP, patient effort, and patient respiratory system mechanics.33

## Slide 250: Auto-Trak Sensitivity
Trilogy automatically adjust its triggering and cycling algorithms to maintain optimum performance in presence of leaks.33
Only available with “passive” circuit configuration.

## Slide 252: DeVilbiss IntelliPAP Bilevel S
Intended for use in treating OSA in spontaneously breathing patients 30 kg (66 lbs) and above by means of application of positive air pressure.34
The device has automatic leak compensation and is intended to be used in home and clinical environments. (pictured in Figure 4-57)

## Slide 253
Figure 4.57

## Slide 255: Continuous Positive Airway Pressure (CPAP)
Pressure-targeted mode of ventilation allowing for spontaneous breathing from a fixed pressure baseline.34
No additional support is provided during inspiration.

## Slide 256: Main Operator Controls Include:
CPAP
Ventilator provides a baseline pressure and the patient breaths spontaneously from that baseline; Vt and flow are variable and will depend on patient effort, and patient respiratory system mechanics.34

## Slide 257: Bilevel
Pressure-targeted mode that delivers patient triggered pressure assist breaths.34

## Slide 258: Main Operator Controls Include:
IPAP
EPAP
Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.34

## Slide 260: DeVilbiss IntelliPAP AutoBilevel
Intended for use in treating OSA in spontaneously breathing patients 30 kg (66 lbs) and above by means of application of positive air pressure.35
Device has automatic leak compensation and is intended to be used in home and clinical environments. (pictured in Figure 4-58)
AutoBilevel is typically used for OSA patients with higher prescribed pressures.
During mode, ventilator automatically increases IPAP in response to clinical events such as hypopnea, apnea, and snoring.
EPAP level stays below the IPAP level by a constant amount.

## Slide 261
Figure 4.58

## Slide 263: Continuous Positive Airway Pressure (CPAP)
Pressure-targeted mode of ventilation allowing for spontaneous breathing from a fixed pressure baseline.35
No additional support is provided during inspiration.

## Slide 264: Main Operator Controls Include:
CPAP
Ventilator provides a baseline pressure and the patient breaths spontaneously from that baseline; Vt and flow are variable and will depend on patient effort, and patient respiratory system mechanics.35

## Slide 265: Bilevel
Pressure-targeted mode that delivers patient triggered pressure assist breaths.35

## Slide 266: Main Operator Controls Include:
IPAP
EPAP
Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.35

## Slide 268: ResMed Lumis Tx
A hospital ventilator designed for patients with sleep disordered breathing, central breathing disorders and respiratory insufficiency.36
It is suitable for a variety of hospital environments.

## Slide 269
Figure 4.59

## Slide 271: CPAP Mode
Lumis delivers a constant baseline level of pressure during inspiration and expiration.36
Patient receives no additional inspiratory pressure during inspiration.

## Slide 272: Main Operator Controls Include:
Set pressure (CPAP)
Vt will be variable and depend on patient effort, and patient respiratory system mechanics.36

## Slide 273: ASV Mode
Pressure-targeted mode in which ventilator automatically adjusts pressure support level to meet a target minute ventilation.36
Designed to consist of solely patient-initiated breaths, but a timed backup mandatory rate does exist that automatically calculates required breath rate to meet target minute ventilation.

## Slide 274: Main Operator Controls Include:
EPAP
Max PS
Min PS
Vt and flow are variable and will depend on the delivered inspiratory pressure level, patient effort, and patient respiratory system mechanics.36

## Slide 275: ASVAuto Mode
ASVAuto mode automatically adjusts expiratory pressure in order to provide only the amount of EPAP required to maintain upper airway patency.36
Designed to consist of solely patient initiated breaths but there does exist a timed backup mandatory rate that automatically calculates required breath rate to meet target minute ventilation.

## Slide 276: Main Operator Controls Include:
Max EPAP
Min EPAP
Max PS
Min PS
Vt and flow are variable and will depend on delivered inspiratory pressure level, patient effort, and patient respiratory system mechanics.36

## Slide 277: Pressure A/C (PAC)
Pressure-targeted mode of ventilation mode that delivers pressure-controlled breaths.36
Inspiration is initiated either by the ventilator at a set rate (time-triggered breath) or by the patient (spontaneous-triggered breath).
Ventilator cycles to expiration when the inspiratory time has been reached.

## Slide 278: Main Operator Controls Include:
IPAP
EPAP
Inspiratory time
RR
Vt will be variable and depend on IPAP, patient effort, and patient respiratory system mechanics.36

## Slide 279: Spontaneous Mode
Bilevel pressure ventilation mode delivering pressure-supported spontaneous breaths.36
Inspiration is initiated only by patient (spontaneous-triggered breath).
Breath cycles to expiration when inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 280: Main Operator Controls Include:
EPAP
IPAP
Cycle
Vt will be variable and depend on IPAP, EPAP, patient effort, and patient respiratory system mechanics.36

## Slide 281: Spontaneous/Timed Mode (S/T)
A bilevel pressure ventilation mode delivering pressure-supported spontaneous breaths.36
Inspiration is initiated either by the ventilator at a set rate (time-triggered breath) or by the patient (spontaneous-triggered breath).
Breath cycles to expiration when inspiratory flow decreases below a preset percentage of inspiratory peak flow.

## Slide 282: Main Operator Controls Include:
EPAP
IPAP
RR
Cycle sensitivity
Vt will be variable and depend on IPAP, EPAP, patient effort, and patient respiratory system mechanics.36

## Slide 283: Timed (T) Mode
Bilevel pressure ventilation mode delivering pressure-targeted breaths.36
Inspiration is initiated only by ventilator at a set rate (time-triggered breath).
Spontaneous breathing not allowed.
Breath cycles to expiration when inspiratory time has been reached.

## Slide 284: Main Operator Controls Include:
EPAP
IPAP
RR
Ti
Vt will be variable and depend on IPAP, EPAP, patient effort, and patient respiratory system mechanics.36

## Slide 285: Intelligent Volume Assured Pressure Support (iVAPS)
Ventilator delivers a variable (from breath to breath) PS level in order to reach a clinician-set target alveolar ventilation.
Ventilator drops the level of support as patient effort increases and conversely increases the PS level when patient effort is too low. iVAPS also has a backup rate, if needed.

## Slide 286: Main Operator Controls Include:
Target Va (L/min)
Target patient rate
Patient height
EPAP
Cycle sensitivity
Vt will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.

## Slide 288: ResMed Astral 100/150
Provides continuous or intermittent ventilatory support for patients weighing more than 11 lbs (5 kg) who require mechanical ventilation.37
The Astral is intended to be used in home, institution/hospital and portable applications for both invasive and noninvasive ventilation and is suitable for a variety of hospital environments. (Figure 4-60)

## Slide 289
Figure 4.60

## Slide 291: (Assisted) Volume-controlled Ventilation (A)CV
Volume-targeted mode of ventilation (VAC) allowing for mandatory, time-triggered breaths or patient-assisted breaths.37
Spontaneous breathing not allowed.
Breath cycles to exhalation once the tidal volume (Vt) has been delivered.

## Slide 292: Main Operator Controls Include:
PEEP
Vt
Inspiratory flow
RR
Inspiratory time
Resulting pressure created is variable and depends on Vt, flow, and patient respiratory system mechanics.37

## Slide 293: Volume-synchronized, Intermittent Mandatory Ventilation (V-SIMV)
Combines two breath types, VAC and either PS or spontaneous breathing, while on CPAP.37
Mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths are all provided.
Ventilator prevents “breath stacking” (a VAC breath compounded by a PS breath) through use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a VAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 294: Main Operator Controls Include:
PEEP
Vt
Flow
SIMV rate
Inspiratory time
Pressure support level
Cycle sensitivity
During VAC breaths, resulting pressure created is variable and depends on Vt, flow, and patient mechanics.37
During PS breaths, volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.37

## Slide 295: Assisted Pressure-controlled Ventilation P(A)CV
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.37
Spontaneous breaths not allowed.
Breath cycles to exhalation once the inspiratory time has been reached.

## Slide 296: Main Operator Controls Include:
PEEP
Inspiratory pressure level
Inspiratory time
RR
Pressure will be consistently delivered from breath to breath; Vt and flow are variable and will depend on inspiratory pressure level, patient effort, and patient respiratory system mechanics.37

## Slide 297: Pressure-synchronized, Intermittent Mandatory Ventilation
Similar to V-SIMV except the target is now pressure instead of volume.37
Combines two breath types, PAC and either PS or spontaneous breathing, while on CPAP.
Combination of mandatory, time-triggered breaths, patient-assisted breaths, and spontaneous breaths.
Ventilator prevents “breath stacking” (a PAC breath compounded by a PS breath) through the use of a “timing window.”
If patient attempts to trigger a breath in timing window, then ventilator will deliver a PAC breath.
If patient attempts to trigger a breath outside window, then a spontaneous breath will occur.

## Slide 298: Main Operator Controls Include:
PEEP
Inspiratory pressure level
Inspiratory time
SIMV rate
Pressure support level
Cycle sensitivity
During PAC breaths, resulting Vt created is variable and depends on inspiratory pressure level, patient effort, and patient respiratory system mechanics.37 During PS breaths, the volume will be variable and depend on the PS level, patient effort, and patient respiratory system mechanics.37

## Slide 299: Pressure Support (PS)
Pressure-targeted mode of ventilation allowing for mandatory, time-triggered breaths or patient-assisted breaths.37
Breath cycles to expiration when inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 300: Main Operator Controls Include:
PEEP
RR
Pressure support level
Cycle sensitivity
Vt will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.37

## Slide 301: Continuous Positive Airway Pressure (CPAP)
A baseline pressure is maintained by the ventilator and then receives no additional support during the spontaneous breaths.37
Other than baseline pressure, patient controls all other aspects of the breath.

## Slide 302: Main Operator Controls Include:
CPAP
The Vt will be variable and depend on patient effort and patient respiratory system mechanics.37

## Slide 303: (Spontaneous) Timed (S/T) Mode (S)T
A bilevel pressure ventilation mode delivering pressure-supported spontaneous breaths.37
Inspiration is initiated either by ventilator at a set rate (time-triggered breath) or by the patient (spontaneous-triggered breath).
Breath cycles to expiration when inspiratory flow decreases below a preset percentage of the inspiratory peak flow.

## Slide 304: Main Operator Controls Include:
EPAP
IPAP
RR
Cycle sensitivity
Vt will be variable and depend on IPAP, EPAP, patient effort, and patient respiratory system mechanics.37

## Slide 305: Pressure (Assist) Control P(A)C
A bilevel pressure ventilation mode delivering pressure-targeted breaths.37
Inspiration is initiated either by ventilator at a set rate (time-triggered breath) or the patient.
Breath cycles to expiration when the inspiratory time has been reached.

## Slide 306: Main Operator Controls Include:
EPAP
IPAP
RR
Ti
Vt will be variable and depend on IPAP, EPAP, patient effort, and patient respiratory system mechanics.37

## Slide 307: Intelligent Volume-assured Pressure Support (iVAPS)
Ventilator delivers a variable (from breath to breath) PS level in order to reach a clinician-set target alveolar ventilation.37
Ventilator drops the level of support as patient effort increases and conversely increases the PS level when patient effort is too low. iVAPS also has a backup rate, if needed.

## Slide 308: Main Operator Controls Include:
Target Va (L/min)
Target patient rate
Patient height
EPAP
Cycle sensitivity
The Vt will be variable and depend on PS level, patient effort, and patient respiratory system mechanics.37